## Sheet1
Standard : IEC62353-Alternative |  | See 3.5 for more details
Equipment :
Test Name |  |  | Value | High Limits | Status
Protective Earth Resistance |  |  | 0.254 Ohm | 0.3 | PASS
Insulation Resistance | Mains to Protective Earth |  | 999 MOhm | - | PASS
| Mains to Applied Parts |  | 999 MOhm | - | PASS
| Applied Parts to Non-Earth Accessible Conductive Part |  | 999 MOhm | - | PASS
Mains Voltage | Live to Neutral |  | 236.1 V | - | PASS
| Neutral to Earth |  | 2.9 V | - | PASS
| Live to Earth |  | 238.6 V | - | PASS
Alternative Equipment Leakage |  | Closed Earth | 557 uAAC | 1000 | PASS
Alternative Equipment Leakage |  | Shorted mains (ecg) | 4.5 uAAC | 50 | PASS